Results of the Quality Assurance Testing

Program for Radiopharmaceuticals 2011

Accredited for compliance

with ISO/IEC 17025

Accreditation Number: 14442

| Prepared by: | Z Ivanov Medical Radiation Services Branch | Signature: | Date: July 2012 |
| --- | --- | --- | --- |
| Approved by: | Ivan Williams A/g Head, Medical Radiation Services Branch | Signature: | Date: July 2012 |

Notice

Australian Radiation Protection and Nuclear Safety Agency 2012

ISSN:  0157-1400

Creative Commons

With the exception of the Commonwealth Coat of Arms, any ARPANSA logos and any content that is marked as being third party material, this publication, Results of the Quality Assurance Testing Program for Radiopharmaceuticals 2011, by the Australian Radiation Protection and Nuclear Safety Agency is licensed under a Creative Commons Attribution 3.0 Australia licence (http://creativecommons.org/licenses/by/3.0/au/). It is a further condition of the licence that any numerical data referred to in this publication may not be changed.

The publication should be attributed as Results of the Quality Assurance Testing Program for Radiopharmaceuticals 2011.

Enquiries regarding the licence and any use of this report are welcome.

ARPANSA
	619 Lower Plenty Road
	YALLAMBIE   VIC   3085

Tel:   1800 022 333 (Freecall) or +61 3 9433 2211
	Email:  info@arpansa.gov.au
	Website:  www.arpansa.gov.au

Disclaimer

All care has been taken in the preparation of this work and its conclusions. However, where the data or results presented are utilised by third parties the Commonwealth of Australia shall not be liable for any special, indirect, consequential or other damages whatsoever resulting from such use. Nor will the Commonwealth of Australia be liable for any damages arising from or in connection with any errors or omissions that have inadvertently occurred in this work.

ABSTRACT

This report tabulates results obtained during 2011 for the Radiopharmaceutical Quality Assurance Test Program conducted by the Australian Radiation Protection and Nuclear Safety Agency (ARPANSA).

CONTENTS

ABSTRACT	iii

INTRODUCTION	1

RESULTS	1

ACKNOWLEDGEMENTS	3

ABBREVIATIONS	4

FLUDEOXYGLUCOSE [18F] INJECTION	5

CHROMIUM [51Cr] EDETATE INJECTION	6

SODIUM CHROMATE [51Cr] SOLUTION	7

GALLIUM [67Ga] CITRATE INJECTION	8

99Mo/ 99mTc CHROMATOGRAPHIC GENERATOR	9

SODIUM PERTECHNETATE [99mTc] INJECTION (FISSION)	10

INDIUM OXINE [111In] SOLUTION	11

IOBENGUANE [123I] INJECTION	12

SODIUM IODIDE [131I] CAPSULES (THERAPY)	13

THALLOUS [201Tl] CHLORIDE INJECTION	14

KIT FOR THE PREPARATION OF TECHNETIUM [99mTc] ALBUMIN AGGREGATED INJECTION (PULMOLITE)	15

KIT FOR THE PREPARATION OF TECHNETIUM [99mTc] BICISATE INJECTION (NEUROLITE)	17

KIT FOR THE PREPARATION OF TECHNETIUM [99mTc] CALCIUM PHYTATE INJECTION (COLLOID)	18

KIT FOR THE PREPARATION OF TECHNETIUM [99mTc] DISOFENIN INJECTION (HEPATOLITE)	19

KIT FOR THE PREPARATION OF TECHNETIUM [99mTc] EXAMETAZIME INJECTION (CERETEC)	20

KIT FOR THE PREPARATION OF TECHNETIUM [99mTc] MEDRONATE INJECTION (MDP)	21

KIT FOR THE PREPARATION OF TECHNETIUM [99mTc] MERTIATIDE INJECTION (MAG3)	22

KIT FOR THE PREPARATION OF TECHNETIUM [99mTc] OXIDRONATE INJECTION (HDP)	23

KIT FOR THE PREPARATION OF TECHNETIUM [99mTc] PENTETATE INJECTION (DTPA)	24

KIT FOR THE PREPARATION OF TECHNETIUM [99mTc] SESTAMIBI INJECTION (CARDIOLITE)	25

KIT FOR THE PREPARATION OF TECHNETIUM [99mTc] SUCCIMER INJECTION (DMSA)	26

KIT FOR THE PREPARATION OF TECHNETIUM [99mTc] TETROFOSMIN (MYOVIEW)	27

KIT FOR THE PREPARATION OF TECHNETIUM [99mTc] TIN PYROPHOSPHATE INJECTION (PYP)	28

INTRODUCTION

The Australian Radiation Protection and Nuclear Safety Agency (ARPANSA) conducts a Radiopharmaceutical Quality Assurance (QA) Test Program under a Memorandum of Understanding (MoU) between ARPANSA and the Therapeutic Goods Administration (TGA). As part of this MoU radiopharmaceuticals used in nuclear medicine in Australia are tested for compliance with specifications. Where the radiopharmaceutical is the subject of a monograph in the British Pharmacopoeia or the European Pharmacopoeia, then the specifications given in these Pharmacopoeias are adopted. Where a monograph is only available in the US Pharmacopoeia, then this specification is generally adopted. Unless stated otherwise, these specifications listed apply at all times up to the date of product expiry. Radionuclidic purity has been determined up to the expiry time.

For Thallous [201Tl] Chloride Injection and Sodium Pertechnetate [99mTc] Injection, radionuclidic purity is determined and the impurity levels at both the time of calibration and expiry are quoted.

Samples for testing were obtained through commercial channels according to the schedule set at the beginning of the test period as described in Quality Assurance of Radiopharmaceuticals including Cold Kits: MR-RPQA-SOP-0000. All technetium-99m cold kits were reconstituted using Sodium Pertechnetate [99mTc] Injection according to the directions supplied in the package insert. Pharmacopoeia methods are used for testing, together with some additional methods described in the ARPANSA Quality System: Radiopharmaceuticals Quality Assurance Testing Program (MR-RPQA-WI-0060A).

RESULTS

The results of testing during 2011 are summarised in the following tables. Overall, 31 batches of 23 different types of radiopharmaceuticals were tested.

In the case of Iobenguane [123I] Injection, the radionuclidic content could not be determined accurately. A certified reference standard for Iodine-123 is still not available and therefore the correction for the effect of the absorption of the abundant low energy X-ray emission by the container and the sample is unknown. Thus, measurement in the original sample glass vial (and using the dose calibrator manufacturer’s setting for 123I) gave an apparent 78 % of the stated radioactivity at the calibration date and time.

For Sodium Iodide [131I] Capsules (Therapy) the BP does not require a “Uniformity of Content” test. The measurement of radioactivity content of 4 capsules of this batch showed that the radioactivity of no capsule differed by more than 7.4 % from the average value, with a relative standard deviation of 7.0 %.

The BP and USP do not require a test for benzyl alcohol used as an antimicrobial preservative in some “ready to use” radiopharmaceuticals. ARPANSA, however, performs this test using an in-house method.

According to the current BP, labelling of the radiopharmaceutical preparations complies with the relevant national and European legislation. In accordance with the “General Requirements for Labels for Medicines” of the Therapeutic Goods Act 1989, TGA Order No. 69, 3(2) (b), (c), the label or labels must include the name(s) and quantity or proportion of all active ingredients in the goods. For small volume injections the label on the primary pack must include the name and quantity of each excipient.

In the case of one type of “cold kit” radiopharmaceutical, the kit from an overseas manufacturer, two batches met the value given in USP (pH = 4.0 – 7.5), which is outside the BP specification (pH = 6.0 – 7.0), but has been approved by the TGA.

Non-compliance of the vial/package label was observed in two batches of ready to use radiopharmaceuticals and one batch of a “cold kit” from an overseas manufacturer. Vial/label non-compliance consisted of the absence of the volume and only the names of the components were listed but not the quantities of each component.

Due to staff and resource reductions, ARPANSA is no longer able to perform animal testing as part of the ARPANSA Quality Assurance Test Program. The Biological Distribution specifications have been retained in the Report for the sake of completeness only.

ACKNOWLEDGEMENTS

The testing of the radiopharmaceuticals and cold kits was performed by Ilonka Bokor and Zlata Ivanov.

ABBREVIATIONS

The following abbreviations are used in the tables –

ARI	-	ANSTO Radiopharmaceuticals and Industrials, Lucas Heights, Sydney, Australia

ANSTO Health	-	ANSTO Australian Nuclear Science and Technology Organisation, Sydney, Australia

Austin Health	-	Heidelberg, VIC. Australia

CIS BIO	-	CIS Bio International, Gif Sur Yvette, France

Cyclotek	-	Bundoora, VIC. Australia

GE Healthcare	-	GE Healthcare Limited, UK

Lantheus (AUST)	-	Lantheus Medical Imaging, Tullamarine, VIC, Australia

Lantheus (USA)	-	Lantheus Medical Imaging, North Billerica, MA, USA

Mall/Covidien	-	Mallinckrodt Medical B.V., Petten, Netherlands, Covidien Pty. Ltd. Lane Cove, NSW, Australia

PHARMA	-	Pharmalucence Inc, Bedford, MA, USA

RADPH	-	Radpharm Scientific, Belconnen, ACT, Australia

RADPH/EDM	-	Edmonton Radiopharmaceutical Centre, Edmonton, Alberta

RADPH/Drax	-	Draxis Specialty Pharmaceuticals Inc., Quebec, Canada

TYCO/MALL	-	TYCO Healthcare, Lane Cove, NSW, Australia; Mallinckrodt Medical B.V., Petten, Netherlands

CALIB. DATE	-	Calibration Date

EXP.	-	Expiry testing

INT.	-	Initial testing

MAX	-	Maximum

MIN	-	Minimum

N.A.	-	Not applicable or not required by BP/USP

N.D.	-	Not detected

No.	-	Number

p	-	Page

reconst.	-	reconstituted

TBD	-	To be done

†	-	Not determined

FLUDEOXYGLUCOSE [18F] INJECTION

Current edition of BP

|  |  | SUPPLIER | Austin Health | Cyclotek |
| --- | --- | --- | --- | --- |
|  |  | LOT/BATCH No. | FDG-111214-1 | 5328 |
|  |  | CALIB. DATE | 14/12/11 @ 9:00h | 14/12/11 @ 6:00h |
|  | SPECIFICATIONS | EXPIRY DATE | 14/12/11 @ 17:00h | 14/12/11 @ 18:00h |
| Appearance | A clear, colourless or slightly yellow yellow solution |  | Pass | Pass |
| Particulate matter | None visible |  | Pass | Pass |
| Identification | The principal peak in the radiochromatogram obtained with the test solution is similar in retention time to the principal peak in the chromatogram obtained with the reference solution. |  | Pass | Pass |
|  | 511 keV peak detected |  | Pass | Pass |
| Radionuclidic content | 90-110% of stated value |  | 105 | 99 |
| Radionuclidic purity | The only gamma photons detected have an energy of 511 keV, a sum peak of 1022 keV may be observed |  | Pass | Pass |
|  | The total radioactivity due to radionuclidic impurities is not more than 0.1 % |  | Pass | Pass |
| pH | 4.5 – 8.5 |  | 6.5 | 6.3 |
| Radiochemical purity (HPLC Method) |  95% of the total radioactivity is present as 2-[18F] fluoro-2 deoxy-D-glucose and 2-[18F] fluoro-2 deoxy-D-mannose | INT EXP | 96.5 ± 0.1 96.5 ± 0.1 | 96.6 ± 0.04 96.4 ± 0.07 |
|  | 2-[18F] fluoro-2 deoxy-mannose: ≤ 10% of the total radioactivity of 2-[18F] fluoro-2 deoxy-glucose and 2-[18F] fluoro-2 deoxy-mannose | INT EXP | 3.5 ± 0.1 4.0 ± 0.2 | 3.4 ± 0.04 3.7 ± 0.07 |
| Vial/Package Label | Supplied |  | Yes | Yes |

CHROMIUM [51Cr] EDETATE INJECTION

Current edition of BP

|  |  | SUPPLIER | ARI | GE Healthcare |
| --- | --- | --- | --- | --- |
|  |  | LOT/BATCH No. | 126352 | 1100 |
|  |  | CALIB. DATE | 01/08/11 | 15/08/11 |
|  | SPECIFICATIONS | EXPIRY DATE | 01/09/11 | 10/10/11 |
| Appearance | A clear, violet solution |  | Pass | Pass |
| Particulate matter | None visible |  | Pass | Pass |
| Identification | Gamma spectrum does not differ significantly from that of a standardised chromium-51 solution |  | Pass | Pass |
| Radionuclidic content | 90-110% of stated value | Vial 1 Vial 2 | 109.9 ± 0.04 109.9 ± 0.05 | 103.3 103.7 |
| Radionuclidic purity | Minimum 99.9 % of the total radioactivity as Chromium-51. The only gamma photons have an energy of 0.320 MeV |  | Pass | Pass |
| pH | 3.5 – 6.5 |  | 5.5 | 4.0 |
| Radiochemical purity |  |  |  |  |
| 1) Chromic ion | as % | INT. 1) | 0.1 ± 0.03 | 0.2 ± 0.1 |
| 2) Chromate ion | as % | 2) | 0.9 ± 0.01 | 0.8 ± 0.2 |
| 3) Cr-edetate | ≥ 95% as 51Cr-edetate | 3) | 99.0 ± 0.03 | 99.0 ± 0.2 |
|  |  | EXP. 1) | 0.2 ± 0.04 | 0.5 ± 0.1 |
|  |  | 2) | 0.4 ± 0.04 | 1.5 ± 0.3 |
|  |  | 3) | 99.4 ± 0.08 | 98.0 ± 0.3 |
| Chromium (Cr) | ≤ 1mg/mL |  | 0.4 | 0.2 |
| Benzyl Alcohol | 90 – 110 % of stated value |  | N.A | 101 |
| Vial/Package Label | Complies |  | Complies | Complies |

SODIUM CHROMATE [51Cr] SOLUTION

Current edition of BP

|  |  | SUPPLIER | GE Healthcare |
| --- | --- | --- | --- |
|  |  | LOT/BATCH No. | 1234 |
|  |  | CALIB. DATE | 10/08/11 |
|  | SPECIFICATIONS | EXPIRY DATE | 05/10/11 |
| Appearance | A clear, colourless or slightly yellow solution |  | Pass |
| Particulate matter | None visible |  | Pass |
| Identification | Gamma spectrum does not differ significantly from that of a standardised chromium-51 solution |  | Pass |
| Radionuclidic content | 90-110% of stated value |  | 105 |
| Radionuclidic purity | Gamma spectrum does not differ significantly from that of a standardised chromium-51 solution |  | Pass |
| pH | 6.0 - 8.5 |  | 6.5 |
| Radiochemical purity | ≥ 90% as sodium chromate | INT. | 99.2 ± 0.03 |
|  | % as chromic ion |  | 0.8 ± 0.03 |
|  |  | EXP. | 99.1 ± 0.02 |
|  |  |  | 0.9 ± 0.02 |
| Total chromate | ≤ 2.7 µg of chromate ion (CrO42-) per MBq |  | 0.3 |
| Benzyl Alcohol | 90 – 110 % of stated value |  | N.A |
| Vial/Package Label | Complies |  | Complies |

GALLIUM [67Ga] CITRATE INJECTION

Current edition of BP

|  |  | SUPPLIER | ANSTO Health | Lantheus (USA) | Mall/Covidien |
| --- | --- | --- | --- | --- | --- |
|  |  | LOT/BATCH No. | 126509-002 | G216111S | 87346 |
|  |  | CALIB. DATE | 09/08/11 | 12/08/11 | 12/08/11 |
|  | SPECIFICATIONS | EXPIRY DATE | 14/08/11 | 19/08/11 | 22/08/11 |
| Appearance | A clear, colourless solution |  | Pass | Pass | Pass |
| Particulate matter | None visible |  | Pass | Pass | Pass |
| Identification | Gamma spectrum does not differ significantly from that of a standardised gallium-67 solution |  | Pass | Pass | Pass |
| Citrate presence | A yellow colour develops in the test solution only |  | Pass | Pass | Pass |
| Radionuclidic content | 90-110% of stated value |  | 102 | 99 | 100 |
| Radionuclidic purity | ≤ 0.2% 66Ga |  | N.D. | N.D. | N.D. |
| pH | 5.0 - 8.0 |  | 7.1 | 6.0 | 7.3 |
| Radiochemical purity | ≥ 97% as Ga Citrate | INT. | 99.5 ± 0.2 | 99.1 ± 0.2 | 98.9 ± 0.2 |
|  |  | EXP. | 99.6 ± 0.1 | 99.2 ± 0.02 | 99.2 ± 0.02 |
| Zinc limit test | ≤ 5 μg/mL |  | < 5 | < 5 | < 5 |
| Benzyl Alcohol | 90 – 110 % of stated value |  | 103 | 96 | † |
| Vial/Package Label | Complies |  | Complies | Complies | Fail |

99Mo/ 99mTc CHROMATOGRAPHIC GENERATOR

Current edition of BP (Sodium Pertechnetate [99mTc] Injection (Fission) and MR-RPQA-WI-0060A

|  |  | SUPPLIER |  | ARI |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | LOT/BATCH No. |  | 125693 - 027 |  |  |  |
|  |  | CALIB. DATE |  | 02/05/11 @ 09:00 h |  |  |  |
|  | SPECIFICATIONS | EXPIRY DATE |  | 16/05/11@ 09:00 h |  |  |  |
| Maximum surface radiation dose rate | < 2000 μSv/h |  |  | 112 μSv/h |  |  |  |
| Dose rate at 1 metre | < 100 μSv/h |  |  | 3.5 μSv/h |  |  |  |
| Appearance (after milking) | A clear, colourless solution |  |  | Pass |  |  |  |
| Particulate matter | None visible |  |  | Pass |  |  |  |
| Identification | Gamma spectrum does not differ significantly from that of a standardised technetium-99m solution |  |  | Pass |  |  |  |
|  |  |  |  | MIN |  | MAX |  |
|  |  |  |  | Elution | Expiry | Elution | Expiry |
| Radionuclidic purity | ≤ 0. 1% 99Mo |  |  | N.D. | N.D. | 1E-4 | 3E-4 |
|  | ≤ 5 x 10-3 % 131I |  |  | N.D. | N.D. | N.D. | N.D. |
|  | ≤ 5 x 10-3 % 103Ru |  |  | N.D. | N.D. | N.D. | N.D. |
|  | ≤ 6 x 10-5 % 89Sr |  |  | † | † | † | † |
|  | ≤ 6 x 10-6 % 90 Sr |  |  | † | † | † | † |
|  | ≤ 1 x 10-7 % alpha-emitting impurities |  |  | † | † | † | † |
|  | ≤ 1 x 10-2 % all other gamma-emitting impurities |  |  | N.D. | N.D. | N.D. | N.D. |
| pH | 4.0 - 8.0 |  |  | 5.0 |  | 5.0 |  |
| Radiochemical purity | ≥ 95% as pertechnetate ion (99mTcO4)- |  | INT. | 99.8 ± 0.02 |  | 99.9 ± 0.01 |  |
|  |  |  | EXP. | 99.9 ± 0.01 |  | 99.9 ± 0.04 |  |
| Aluminium | ≤ 5μg/mL |  |  | 0.1 μg/mL |  | 0.1 μg/mL |  |
| Milking efficiency | None (for information only) |  |  | 90.0 % |  | 91.5 % |  |
| Moly assay | ≤ 0.1 % 99Mo at expiry |  |  | N.D. |  | 3 E-4 |  |
| (99Mo breakthrough) |  |  |  |  |  |  |  |
| Vial/Package Label | Supplied |  |  | Supplied |  |  |  |

SODIUM PERTECHNETATE [99mTc] INJECTION (FISSION)

Current edition of BP

|  |  | SUPPLIER | Lantheus (AUS) |  |
| --- | --- | --- | --- | --- |
|  |  | LOT/BATCH No. | 422757 |  |
|  |  | CALIB. DATE | 24/11/11 @ 09:00 h |  |
|  | SPECIFICATIONS | EXPIRY DATE | 24/11/11 @ 17:00 h |  |
| Appearance | A clear, colourless solution |  | Pass |  |
| Particulate matter | None visible |  | Pass |  |
| Identification | Gamma spectrum does not differ significantly from that of a standardised technetium-99m solution |  | Pass |  |
| Radionuclidic content | 90-110% of stated value |  | 100 |  |
|  |  |  | INT. | EXP. |
| Radionuclidic purity | ≤ 0. 1% 99Mo |  | 3E-4 | 7E-4 |
|  | ≤ 5 x 10-3 % 131I |  | 2E-5 | 5E-5 |
|  | ≤ 5 x 10-3 % 103Ru |  | N.D. | N.D. |
|  | ≤ 6 x 10-5 % 89Sr |  | † | † |
|  | ≤ 6 x 10-6 % 90 Sr |  | † | † |
|  | ≤ 1 x 10-7 % alpha-emitting impurities |  | † | † |
|  | ≤ 1 x 10-2 % all other gamma-emitting impurities |  | N.D. | N.D. |
| pH | 4.0 - 8.0 |  | 4.5 |  |
| Radiochemical purity | ≥ 95% as pertechnetate ion (99mTcO4-) |  | 99.9 ± 0.01 | 99.9 ± 0.02 |
| Aluminium (Al) | ≤ 5 μg /mL |  | †* |  |
| Vial/Package Label | Complies |  | N.A. |  |

*Insufficient sample size.

INDIUM OXINE [111In] SOLUTION

Current edition of BP

|  |  | SUPPLIER | GE Healthcare |
| --- | --- | --- | --- |
|  |  | LOT/BATCH No. | 6178 |
|  |  | CALIB. DATE | 12/10/11 |
|  | SPECIFICATIONS | EXPIRY DATE | 17/10/11 |
| Appearance | A clear, colourless solution |  | Pass |
| Particulate matter | None visible |  | Pass |
| Identification | Gamma spectrum does not differ significantly from that of a standardised indium-111 solution, apart from any difference due to the presence of indium-114m |  | Pass |
| Radionuclidic content | 90-110% of stated value |  | 110 |
| Radionuclidic purity | Gamma spectrum does not differ significantly from that of a standardised indium-111 solution, apart from any difference due to the presence of indium-114m |  | Pass |
|  | ≤ 0.25% of the total radioactivity is due to radionuclides other than indium-111 at all times up to expiry |  | N.D. |
| pH | 6.0 – 7.5 |  | 7.0 |
| Radiochemical purity | ≥ 90% of activity as 111In-Oxine | INT. | 94.8 ± 1.2 |
|  |  | EXP. | 93.7 ± 0.5 |
| Vial/Package Label | Complies |  | Complies |

IOBENGUANE [123I] INJECTION

Current edition of BP

|  |  | SUPPLIER | ANSTO Health |  |
| --- | --- | --- | --- | --- |
|  |  | LOT/BATCH No. | 127093-012 |  |
|  |  | CALIB. DATE | 18/10/11 @ 09:00 h |  |
|  | SPECIFICATIONS | EXPIRY TIME | 19/10/11 @ 01:00 h |  |
| Appearance | A clear, colourless or slightly yellow solution |  | Pass |  |
| Particulate matter | None visible |  | Pass |  |
| Identification | Gamma spectrum does not differ significantly from that of a standardised iodine-123 solution apart from any differences attributed to the presence of iodine-125, tellutium-121 and other radionuclidic impurities |  | Pass |  |
| Radionuclidic content | 90-110% of stated value |  | 78* |  |
| Radionuclidic purity | ≤ 0.35 % of the total radioactivity is due to radionuclides other than iodine-123 |  | Pass |  |
| pH | 3.5 – 8.0 |  | 5.0 |  |
|  |  |  | INT. | EXP. |
| Radiochemical purity | ≥ 95% of activity as iobenguane |  | 98.6 ± 0.02 | 98.84 ± 0.1 |
|  | ≤ 4 % of activity as iodide |  | 0.9 ± 0.1 | 0.83 ± 0.04 |
|  | ≤ 1% of activity in other peaks |  | 0.5 ± 0.1 | 0.33 ± 0.02 |
| Vial/Package Label | Complies |  | Complies |  |

* Measurements performed in the glass vial supplied by the manufacturer.

Note: No reference standard is available and no adjustment was made to the Capintec setting. For details refer to p 5 of this report.

SODIUM IODIDE [131I] CAPSULES (THERAPY)

Current edition of BP

|  |  | SUPPLIER | ANSTO Health |
| --- | --- | --- | --- |
|  |  | LOT/BATCH No. | 127070-026 |
|  |  | CALIB. DATE | 24/10/11 |
|  | SPECIFICATIONS | EXPIRY DATE | 07/11/11 |
| Appearance | Gelatine capsule |  | Pass |
| Identification | Gamma spectrum does not differ significantly from that of a standardised iodine-131 solution |  | Pass |
| Radionuclidic content | 90-110% of stated value |  | 101 ± 7* |
| Radionuclidic purity |  99.9% as 131I, |  | Pass |
|  |  0.1% of the total radioactivity is due to 130I, 133I, 135I and other radionuclidic impurities |  | Pass |
| Radiochemical purity |  95% of activity as iodide | INIT. | 100 |
|  |  | EXP. | 100 |
| Disintegration | The shell and its contents dissolve completely within 15 min. |  | Pass |
| Vial/Package Label | Complies |  | Complies |

* 4 capsules measured.

Note: The BP does not require a Uniformity of Content test for Sodium Iodide [131I] Capsules (Therapy). For details refer to p 6 of this report.

THALLOUS [201Tl] CHLORIDE INJECTION

Current edition of BP

|  |  |  | SUPPLIER | ARI |  | Mall/Covidien |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | LOT/BATCH No. | 125477-002 |  | 85731 |  |
|  |  |  | CALIB. DATE | 31/03/11 |  | 31/03/11 |  |
|  | SPECIFICATIONS |  | EXPIRY DATE | 05/04/11 |  | 07/04/11 |  |
| Appearance | A clear colourless solution |  |  | Pass |  | Pass |  |
| Particulate matter | None visible |  |  | Pass |  | Pass |  |
| Identification | Gamma spectrum does not differ significantly from that of a standardised thallium-201 solution |  |  | Pass |  | Pass |  |
| Radionuclidic content | 90-110% of stated value |  |  | 105 ± 0.06* |  | 102 ± 0.01* |  |
| Radionuclidic purity | At all times up to expiry |  |  | At calibration | At expiry | At calibration | At expiry |
|  | 201Tl | ≥ 97 % |  | 99.6 | 99.1 | 99.6 | 98.9 |
|  | 202Tl | ≤ 2.0 % |  | 0.37 | 0.87 | 0.3 | 1.05 |
|  | 200Tl | % |  | 0.04 | 0.005 | 0.08 | 0.004 |
|  | 201Pb | % |  | N.D. | N.D. | N.D. | N.D. |
|  | 203Pb | % |  | N.D. | N.D. | N.D. | N.D. |
| pH | 4.0 - 7.0 |  |  | 4.5 |  | 5.0 |  |
| Radiochemical purity | ≥ 95% of the activity is present as Thallous ion |  | INT. | 98.6 ± 0.3 |  | 99.5 ± 0.04 |  |
|  |  |  | EXP. | 99.1 ± 0.2 |  | 99.2 ± 0.03 |  |
| Thallium | ≤ 10 μg/mL |  |  | < 2.5 μg/mL |  | < 1 μg/mL |  |
| Benzyl Alcohol | 90 – 120 % of stated value** |  |  | 118 |  | N.A. |  |
| Vial/Package Label | Complies |  |  | Complies |  | Fail |  |

* Measurements performed in two different dose calibrators (Capintec ARC-120 and ARC-15R).

** Manufacturer’s approved specification.

KIT FOR THE PREPARATION OF TECHNETIUM [99mTc] ALBUMIN AGGREGATED INJECTION (PULMOLITE)

Current edition of USP/BP

|  |  | SUPPLIER | PHARMA |  | RADPH/EDM |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | LOT/BATCH No. | 160038A |  | 9D399 |  |
|  | SPECIFICATIONS | EXPIRY DATE | 28/02/11 |  | Sept 2011 |  |
|  |  |  | INT. | EXP. | INT. | EXP. |
| Appearance before reconstitution | Freeze-dried solid with no evidence of moisture |  | Pass | Pass | Pass | Pass |
| Appearance after reconstitution | A white suspension which may separate on standing* |  | Pass | Pass | Pass | Pass |
| Check for vacuum | Vacuum present. If no vacuum is found, the vial should be discarded. |  | N.A | N.A | N.A. | N.A. |
| pH | 3.8 - 7.5*after reconstitution 3.8 - 8.0**after reconstitution |  | 6.0 |  | 5.0 |  |
| Radiochemical purity | 1) ≥ 90.0 % in aggregated albumin (by chromatography) |  | 99.9 ± 0.01 | 98.6 ± 0.1 | 99.2 ± 0.1 | 99.2 ± 0.1 |
|  | 2) ≤ 10 % as soluble and dispersed radiochemical impurities (by centrifugation) |  | 2.3 ± 0.05 | 6.5 ± 0.1 | 1.6 ± 0.04 | 0.8 ± 0.04 |
| Particle size | ≥ 90% of the observed aggregated particles (not less than 100) have a diameter between 10 μm and 90 μm |  | Pass | Pass | Pass | Pass |
|  | No particle having a maximum diameter > 150 μm is present |  | Pass | Pass | Pass | Pass |
| Non filterable radioactivity* | The radioactivity remaining on the membrane is ≥ 90% |  | 97.8 ± 0.5 | 97.3 ± 0.4 | 98.5 ± 0.02 | 99.0 ± 0.2 |
| Tin content* | ≤ 3 mg/mL of Sn |  | < 3 |  | < 3 |  |
| Biological distribution | ≥ 80% in the lungs |  | † |  | † |  |
|  | ≤ 5% in the liver + spleen |  | † |  | † |  |
| Vial/Package Label | Complies |  | Complies |  | Complies |  |

* Current edition of BP.

** Current edition of USP

KIT FOR THE PREPARATION OF TECHNETIUM [99mTc] ALBUMIN AGGREGATED INJECTION (PULMOLITE)

Current edition of USP/BP

|  |  | SUPPLIER | RADPH/Drax |  |
| --- | --- | --- | --- | --- |
|  |  | LOT/BATCH No. | QC261 |  |
|  | SPECIFICATIONS | EXPIRY DATE | Sept.2012 |  |
|  |  |  | INT. | EXP. |
| Appearance before reconstitution | Freeze-dried solid with no evidence of moisture |  | Pass | TBD |
| Appearance after reconstitution | A white suspension which may separate on standing* |  | Pass | TBD |
| Check for vacuum | Vacuum present. If no vacuum is found, the vial should be discarded. |  | N.A | TBD |
| pH | 3.8 - 7.5*after reconstitution 3.8 - 8.0**after reconstitution |  | 5.5 | TBD |
| Radiochemical purity | 1) ≥ 90.0 % in aggregated albumin (by chromatography) |  | 99.3 ± 0.03 | TBD |
|  | 2) ≤ 10 % as soluble and dispersed radiochemical impurities (by centrifugation) |  | 1.0 ± 0.1 | TBD |
| Particle size | ≥ 90% of the observed aggregated particles (not less than 100) have a diameter between 10 μm and 90 μm |  | Pass | TBD |
|  | No particle having a maximum diameter > 150 μm is present |  | Pass | TBD |
| Non filterable radioactivity* | The radioactivity remaining on the membrane is ≥ 90% |  | 99.2 ± 0.05 |  |
| Tin content* | ≤ 3 mg/mL of Sn |  | TBD |  |
| Biological distribution | ≥ 80% in the lungs |  | † |  |
|  | ≤ 5% in the liver + spleen |  | † |  |
| Vial/Package Label | Complies |  | Complies |  |

* Current edition of BP.

** Current edition of USP

KIT FOR THE PREPARATION OF TECHNETIUM [99mTc] BICISATE INJECTION (NEUROLITE)

Current edition of BP/USP

|  |  | SUPPLIER | Lantheus (USA) |  |
| --- | --- | --- | --- | --- |
|  |  | LOT/BATCH No. | 0200A |  |
|  | SPECIFICATIONS | EXPIRY DATE | 01/10/11 |  |
|  |  |  | INT. | EXP. |
| Appearance before reconstitution | Freeze-dried solid |  | Pass | Pass |
| Appearance after reconstitution | A clear, colourless solution |  | Pass | Pass |
| pH | 6.5 - 7.5* after reconstitution |  | 7.0 | 7.0 |
| Radiochemical purity | ≥ 94.0 % as 99mTc-Bicisate* ≥ 90.0 % as 99mTc-Bicisate** |  | 95.6 ± 0.1 | 97.2 ± 0.1 |
|  | ≤ 6.0 % of the total radioactivity as sum of impurities* ≤ 10.0 % as impurities** |  | 4.4 ± 0.1 | 2.8 ± 0.1 |
| Tin content | 12 - 72 μg SnCl2.2H2O*** |  | N.A. |  |
| Vial/Package Label | Complies |  | Complies |  |

* Current edition of BP specification.

** Current edition of USP specification.

***Value given in label/product information.

KIT FOR THE PREPARATION OF TECHNETIUM [99mTc] CALCIUM PHYTATE INJECTION (COLLOID)

Current edition of BP (Technetium [99mTc] Colloidal Tin Injection)

|  |  | SUPPLIER | RADPH |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | LOT/BATCH No. | 2823/2879 |  |  |  |
|  | SPECIFICATIONS | EXPIRY DATE | January 2012 |  |  |  |
|  |  |  | INT. |  | EXP. |  |
| Appearance before reconstitution | Freeze-dried solid with no evidence of moisture |  | Pass |  | Pass |  |
| Appearance after reconstitution | A clear, colourless or slightly yellow solution, free of any visible particulate matter |  | Pass |  | Pass |  |
| Check for vacuum | If a vacuum is not observed when the vial is pierced, the vial should be discarded |  | Pass |  | Pass |  |
| pH | 4.0 - 7.0 after reconstitution |  | 4.5 |  | 4.0 |  |
|  |  |  | Pall ITLC-SG | Varian ITLC-SG | Pall ITLC-SG | Varian ITLC-SG |
| Radiochemical purity | ≥ 95.0 % as 99mTc-colloid |  | 97.1 ± 0.6 | 97.5 ± 0.2 | 98.9 ± 0.6 | 98.5 ± 0.2 |
| Tin content | ≤ 1.0 mg SnCl2 |  | 0.4 |  |  |  |
| Vial/Package Label | Complies |  | Complies |  |  |  |

KIT FOR THE PREPARATION OF TECHNETIUM [99mTc] DISOFENIN INJECTION (HEPATOLITE)

Current edition of USP

|  |  | SUPPLIER | PHARMA |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | LOT/BATCH No. | 230027A |  |  |  |
|  | SPECIFICATIONS | EXPIRY DATE | 30/09/12 |  |  |  |
|  |  |  | INT. |  | EXP. |  |
| Appearance before reconstitution | Freeze-dried solid with no evidence of moisture |  | Pass |  | Pass |  |
| Appearance after reconstitution | A clear, colourless or slightly yellow solution, free of any visible particulate matter |  | Pass |  | Pass |  |
| pH | 4.0 - 5.0 after reconstitution |  | 4.5 |  | 4.4 |  |
|  |  |  | Pall ITLC-SG | Varian ITLC-SG | Pall ITLC-SG | Varian ITLC-SG |
| Radiochemical purity | 1) ≥ 90.0 % as 99mTc-disofenin | 1) | 99.4 |  | 99.3 |  |
|  | 2) hydrolysed 99mTc + tin colloid (chromatography system A) | 2) | 0.2 ± 0.07 | 0.4 ± 0.05 | 0.2 ± 0.02 | 0.3 ± 0.03 |
|  | 3) free pertechnetate (chromatography system B) | 3) | 0.4 ± 0.2 |  | 0.5 ± 0.11 |  |
|  | 2) +3 ) ≤10% |  | 0.6 |  | 0.7 |  |
| Tin content | 0.24 - 0.6 mg SnCl2* |  | N.A. |  | N.A. |  |
| Biological distribution |  70% in gallbladder + intestines |  | † |  | † |  |
|  |  10% in the liver |  | † |  | † |  |
|  | ≤ 10% in the kidneys |  | † |  | † |  |
|  | ≤ 3 % in the stomach |  | † |  | † |  |
|  | ≤ 3 % in the blood |  | † |  | † |  |
| Vial/Package Label | Complies |  | Complies |  | Complies |  |

* Value given in label/product information.

KIT FOR THE PREPARATION OF TECHNETIUM [99mTc] EXAMETAZIME INJECTION (CERETEC)

Current edition of BP

|  |  | SUPPLIER | GE Healthcare |  |
| --- | --- | --- | --- | --- |
|  |  | LOT/BATCH No. | 1342 |  |
|  | SPECIFICATIONS | EXPIRY | 18/11/11 |  |
|  |  |  | INT. | EXP. |
| Appearance before reconstitution | Freeze-dried solid with no evidence of moisture |  | Pass | Pass |
| Appearance after reconstitution | A clear, colourless solution, free of any visible particulate matter |  | Pass | Pass |
| pH | 9.0 – 9.8 after reconstitution |  | 9.1 | 9.1 |
| Radiochemical purity | ≥ 80 % total radioactivity of: lipophilic 99mTc-Exametazime (X) and meso isomer of lipophilic 99mTc-Exametazime (A) (X+A)=100 – (C+B+D+E) |  | 94.5 | 89.8 |
|  |  10 % as free pertechnetate (C) |  | 1.8 ± 0.2 | 6.0 ± 0.6 |
|  | % sum of: (B+D+E) 99mTc-colloidal (B), non lipophilic 99mTc-Exametazime complex (D) and meso isomer of non lipophilic 99mTc-Exametazime complex (E) |  | 3.7 ± 0.2 | 4.2 ± 0.4 |
|  |  5% of the radioactivity due to: meso isomer of lipophilic 99mTc-Exametazime (A) |  | 0.1 ± 0.1 | 0.3 ± 0.1 |
| Tin content | 7.6 μg SnCl2.2H2O/vial* |  | N.A. |  |
| Biological distribution |  1.5 % in the brain |  | † |  |
|  |  20 % in the intestines |  | † |  |
|  |  15 % in the liver |  | † |  |
| Vial/Package Label | Complies |  | Complies |  |

* Value given in label/product information.

KIT FOR THE PREPARATION OF TECHNETIUM [99mTc] MEDRONATE INJECTION (MDP)

Current edition of BP

|  |  | SUPPLIER | ANSTO Health |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | LOT/BATCH No. | 2920 |  |  |  |
|  | SPECIFICATIONS | EXPIRY DATE | March 2012 |  |  |  |
|  |  |  | INT. |  | EXP. |  |
| Appearance before reconstitution | Freeze-dried solid with no evidence of moisture |  | Pass |  | Pass |  |
| Appearance after reconstitution | A clear, colourless solution |  | Pass |  | Pass |  |
| pH | 3.5 – 7.5 after reconstitution |  | 4.5 |  | 4.5 |  |
|  |  |  | Pall ITLC-SG | Varian ITLC-SG | Pall ITLC-SG | Varian ITLC-SG |
| Radiochemical purity | 1)  95.0 % as 99mTc-MDP |  | 99.8 | 99.7 |  | 98.8 |
|  | 2)  2.0 % as 99mTcO4- |  | 0.1 ± 0.05 | 0.1 ± 0.03 | † | 0.2 ± 0.05 |
|  | 3) as colloidal 99mTc |  | 0.1 ± 0.01 | 0.2 ± 0.1 | † | 1.0 ± 0.3 |
|  | 2) + 3)  5.0 % |  | 0.2 | 0.3 |  | 1.2 |
| Tin content |  3 mg/mL |  |  |  |  |  |
|  | 0.84 mg SnCl2* |  | 0.81 mg |  |  |  |
| Biological distribution |  1.5% attached to femur |  | † |  | † |  |
|  |  1.0% in the liver |  | † |  | † |  |
|  |  0.05 %/g in the blood |  | † |  | † |  |
| Vial/Package Label | Complies |  | Complies |  | Complies |  |

* Value given in label/product information.

KIT FOR THE PREPARATION OF TECHNETIUM [99mTc] MERTIATIDE INJECTION (MAG3)

Current edition of BP

|  |  |  |  | SUPPLIER |  | TYCO/MALL |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | LOT/BATCH No. |  | 09690013 |  |  |  |
|  |  | SPECIFICATIONS |  | EXPIRY DATE |  | 16/09/11 |  |  |  |
|  |  |  |  |  |  | INT. |  | EXP. |  |
| Appearance before reconstitution |  | Freeze-dried solid with no evidence of moisture |  |  |  | Pass |  |  |  |
| Appearance after reconstitution |  | A clear, colourless solution |  |  |  | Pass |  |  |  |
| pH |  | 5.0 - 7.5 after reconstitution |  |  |  | 5.7 |  | 5.6 |  |
| Radiochemical purity |  | ≥ 94.0 % as 99mTc-MAG3 |  |  |  | 97.9 ± 0.1* | 99.2 ± 0.04** | 98.3 ± 0.2* | 99.3 ± 0.2** |
|  |  | ≤ 3.0 % as hydrophilic impurities including [99mTc] pertechnetate |  |  |  | 1.7 ± 0.2* |  | 1.5± 0.3* |  |
|  |  | ≤ 4.0 % as lipophilic impurities |  |  |  | 0.4 ± 0.03* |  | 0.1 ± 0.07* |  |
|  |  | ≤ 2% as reduced-hydrolysed technetium (by chromatography) |  |  |  | 0.03 ± 0.005* |  | 0.1 ± 0.01* |  |
| Tin content |  | ≥ 50 μg SnCl2.2H2O/vial*** |  |  |  | N.A. |  |  |  |
| Vial/Package Label | Complies |  |  |  | Complies |  |  |  |  |

* Tested by Current addition of BP

**Tested by the method recommended by the manufacturer (Sep-Pak C18 cartridge).

*** Value given in label/product information.

KIT FOR THE PREPARATION OF TECHNETIUM [99mTc] OXIDRONATE INJECTION (HDP)

Current edition of USP

|  |  | SUPPLIER | TYCO/MALL |  |
| --- | --- | --- | --- | --- |
|  |  | LOT/BATCH No. | 0911007 |  |
|  | SPECIFICATIONS | EXPIRY DATE | 11/12/11 |  |
|  |  |  | INT. | EXP. |
| Appearance before reconstitution | Freeze-dried solid with no evidence of moisture |  | Pass | Pass |
| Appearance after reconstitution | A clear, colourless solution |  | Pass | Pass |
| pH | 2.5 – 7.0 after reconstitution |  | 4.5 | 4.5 |
| Radiochemical purity* | 1)  90.0 % as 99mTc-oxidronate |  | 97.8 | 97.1 |
|  | 2) % as free pertechnetate |  | 1.9 ± 0.6 | 2.5 ± 0.2 |
|  | 3) % as colloidal 99mTc |  | 0.3 ± 0.04 | 0.4 ± 0.03 |
|  | 2) + 3) ≤ 10 % |  | 2.2 | 2.9 |
| Tin content | 0.258 - 0.342 mg SnCl2.2H2O** |  | N/A |  |
| Biological distribution |  1.0% attached to one femur |  | † |  |
|  |  5.0% in the liver |  | † |  |
|  |  5.0% in the kidneys |  | † |  |
| Vial/Package Label | Complies |  | Complies |  |

*Value given in label/product information.

KIT FOR THE PREPARATION OF TECHNETIUM [99mTc] PENTETATE INJECTION (DTPA)

Current edition of BP

|  |  | SUPPLIER | RADPH |  |  |  |  |  | ARI |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | LOT/BATCH No. | 2800 |  |  |  |  |  | 2852 |  |  |  |  |
|  | SPECIFICATIONS |  | July 2011 |  |  |  |  |  | August 2011 |  |  |  |  |
|  |  |  | INT. |  |  | EXP. |  |  | INT. |  |  | EXP. |  |
| Appearance before reconstitution | Freeze-dried solid with no evidence of moisture |  | Pass |  |  | Pass |  |  | Pass |  |  | Pass |  |
| Appearance after reconstitution | A clear, colourless or slightly yellow solution |  | Pass |  |  | Pass |  |  | Pass |  |  | Pass |  |
| Check for vacuum | If a vacuum is not observed when the vial is pierced, the vial should be discarded |  | Pass |  |  | Pass |  |  | Pass |  |  | Pass |  |
| pH | 4.0 – 7.5 after reconstitution |  | 5.5 |  |  | 5.7 |  |  | 4.2 |  |  | 4.2 |  |
|  |  |  | Pall ITLC-SG | Varian ITLC-SG | Pall ITLC-SG |  | Varian ITLC-SG | Pall ITLC-SG |  | Varian ITLC-SG | Pall ITLC-SG |  | Varian ITLC-SG |
| Radiochemical purity | 1)  95.0 % as 99mTc-DTPA |  | 99.9 | 99.8 | 99.6 |  | 99.6 | 99.9 |  | 99.3 | 99.8 |  | 99.7 |
|  | 2) Colloidal 99mTc impurity (chromatography system A) |  | 0.06 ± 0.02 | 0.1 ± 0.02 | 0.3 ± 0.01 |  | 0.3 ± 0.02 | 0.07 ± 0.01 |  | 0.4 ± 0.1 | 0.09 ± 0.01 |  | 0.11 ± 0.03 |
|  | 3) Free pertechnetate 99mTc (chromatography system B) |  | 0.04 ± 0.02 | 0.1 ± 0.07 | 0.1 ± 0.02 |  | 0.1 ± 0.01 | 0.06 ± 0.02 |  | 0.31 ± 0.06 | 0.08 ± 0.01 |  | 0.15 ± 0.02 |
|  | 2 ) + 3) ≤ 5.0% |  | 0.1 | 0.2 | 0.4 |  | 0.4 | 0.1 |  | 0.7 | 0.2 |  | 0.3 |
| Tin content | ≤ 1 mg/mL |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 0.21 mg SnCl2 /vial* |  |  |  |  |  |  | 0.18 ± 0.02mg |  |  |  |  |  |
|  | 1 mg SnCl2 /vial* |  | 0.4 ± 0.01mg |  |  |  |  |  |  |  |  |  |  |
| Vial/Package Label | Complies |  | Complies |  |  |  |  | Complies |  |  |  |  |  |

* Value given in label/product information.

KIT FOR THE PREPARATION OF TECHNETIUM [99mTc] SESTAMIBI INJECTION (CARDIOLITE)

Current edition of BP

|  |  | SUPPLIER | Lantheus (USA) |  | Lantheus (USA) |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | LOT/BATCH No. | 4009 |  | 4027 |  |
|  | SPECIFICATIONS | EXPIRY DATE | 01/03/12 |  | 01/06/12 |  |
|  |  |  | INT. | EXP. | INT. | EXP. |
| Appearance before reconstitution | Freeze dried solid with no evidence of moisture |  | Pass | Pass | Pass | † |
| Appearance after reconstitution | A clear, colourless solution |  | Pass | Pass | Pass | † |
| pH | 5.0 - 6.0 after reconstitution |  | 5.2 |  | 5.2 |  |
| Radiochemical purity |  94.0 % as 99mTc-Sestamibi |  | 98.8 ± 0.01 | 98.7± 0.03 | 98.5 ± 0.09 | † |
|  |  |  | 1.2 ± 0.01 | 1.3 ± 0.03 | 1.5 ± 0.09 | † |
| Tin content | 0.075 mg SnCl2* |  | N.A. |  | N.A. |  |
| Vial/Package Label | Complies |  | Complies |  | Complies |  |

* Value given in label/product information.

KIT FOR THE PREPARATION OF TECHNETIUM [99mTc] SUCCIMER INJECTION (DMSA)

Current edition of BP

|  |  | SUPPLIER | CIS-BIO |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | LOT/BATCH No. | A004G |  |  |  |
|  | SPECIFICATIONS | EXPIRY DATE | 30/07/11 |  |  |  |
|  |  |  | INT. |  | EXP. |  |
| Appearance before reconstitution | Freeze dried solid |  | Pass |  | Pass |  |
| Appearance after reconstitution | A clear, colourless solution |  | Pass |  | Pass |  |
| Check for vacuum | If a vacuum is not observed when the vial is pierced, the vial should be discarded |  | Pass |  | Pass |  |
| pH | 2.3 - 3.5 after reconstitution |  | 2.8 |  | 2.9 |  |
|  |  |  | Pall ITLC-SG | Varian ITLC-SG | Pall ITLC-SG | Varian ITLC-SG |
| Radiochemical purity |  95.0 % as 99mTc-DMSA |  | 99.93 | 99.95 | 99.9 | 99.9 |
|  |  2.0 % as 99mTcO4- |  | 0.07 ± 0.04 | 0.05 ± 0.01 | 0.1 ± 0.03 | 0.1 ± 0.04 |
| Tin content |  1 mg/mL |  | < 1 |  |  |  |
| Biological distribution |  40% in the kidneys |  | † |  | † |  |
|  |  10% in the liver |  | † |  | † |  |
|  |  2% in the stomach |  | † |  | † |  |
|  |  5% in the lungs |  | † |  | † |  |
| Vial/Package Label | Complies |  | Fail |  |  |  |

KIT FOR THE PREPARATION OF TECHNETIUM [99mTc] TETROFOSMIN (MYOVIEW)

Current edition of USP

|  |  | SUPPLIER | GE Healthcare |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | LOT/BATCH No. | 1738 |  |  |  |
|  | SPECIFICATIONS | EXPIRY DATE | 11/11/11 |  |  |  |
|  |  |  | INT. |  | EXP. |  |
| Appearance before reconstitution | Freeze-dried solid with no evidence of moisture |  | Pass |  | Pass |  |
| Appearance after reconstitution | A clear, colourless solution, free of any visible particulate matter |  | Pass |  | Pass |  |
| pH | 8.3 - 9.1 after reconstitution |  | 8.3 |  | 8.3 |  |
|  |  |  | Pall ITLC-SG | Varian ITLC-SG | Pall ITLC-SG | Varian ITLC-SG |
| Radiochemical purity | 1)  90.0 % as 99mTc-Tetrofosmin |  | 97.1 ± 0.1 | 97.5 ± 0.2 | 98.0 * | 98.5 ± 0.05 |
|  | 2) Reduced hydrolysed & hydrophilic impurities |  | 2.3 ± 0.02 | 2.2 ± 0.1 | 1.6 * | 1.3 ± 0.03 |
|  | 3) Unbound pertechnetate |  | 0.5 ± 0.1 | 0.3 ± 0.02 | 0.4 * | 0.2 ± 0.02 |
|  | 2) + 3) ≤ 10% |  | 2.8 | 2.5 | 2.0 | 1.5 |
| Tin content | 0.03 mg/vial SnCl2.2H2O** |  | N.A. |  |  |  |
| Vial/Package Label | Complies |  | Complies |  |  |  |

* Single strip only.

** Value given in label/product information.

KIT FOR THE PREPARATION OF TECHNETIUM [99mTc] TIN PYROPHOSPHATE INJECTION (PYP)

Current edition of BP/USP

|  |  | SUPPLIER | Covidien/Mall |  |  |  | CIS BIO |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | LOT/BATCH No. | 0941002 |  |  |  | A006D |  |  |  |
|  | SPECIFICATIONS | EXPIRY DATE | 07/09/11 |  |  |  | 28/10/11 |  |  |  |
|  |  |  | INT. |  | EXP. |  | INT. |  | EXP. |  |
| Appearance before reconstitution | Freeze-dried solid with no evidence of moisture |  | Pass |  |  |  | Pass |  | Pass |  |
| Appearance after reconstitution | A clear, colourless or slightly yellow solution, free of any visible particulate matter |  | Pass |  |  |  | Pass |  | Pass |  |
| Check for vacuum | If a vacuum is not observed when the vial is pierced, the vial should be discarded |  | N.A. |  |  |  | N.A. |  | N.A. |  |
| pH | 6.0 – 7.0 after reconstitution* |  |  |  |  |  |  |  |  |  |
|  | 4.0 - 7.5 after reconstitution** |  | 5.0 |  | 4.5 |  | 5.0 |  | 5.0 |  |
|  |  |  | Pall ITLC-SG | Varian ITLC-SG | Pall ITLC-SG | Varian ITLC-SG | Pall ITLC-SG | Varian ITLC-SG | Pall ITLC-SG | Varian ITLC-SG |
| Radiochemical purity | 1)  90.0 % as 99mTc-PYP |  | 98.8 | 95.0 | 98.6 | 93.8 | 99.0 | 91.3 | 98.9 | 93.0 |
|  | 2) as free pertechnetate |  | 0.4 ± 0.2 | 2.7 ± 0.9 | 0.8 ± 0.2 | 4.5 ± 0.8 | 0.3 ± 0.02 | 4.5 ± 0.7 | 0.3 ± 0.2 | 4.7 ± 0.5 |
|  | 3) as colloidal [99mTc] |  | 0.8 ± 0.1 | 2.3 ± 0.1 | 0.6 ± 0.2 | 1.6 ± 0.2 | 0.7 ± 0.2 | 4.7 ± 0.8 | 0.8 ± 0.1 | 2.3± 0.6 |
|  | 2) + 3) ≤ 10 % |  | 1.2 | 5.0 | 1.4 | 6.2 | 1.0 | 8.7 | 1.1 | 7.0 |
| Sodium pyrophosphate | 1-50 mg/mL sodium pyrophosphate on reconstitution |  | 12 mg/mL |  |  |  | 20 mg/mL |  |  |  |
| Tin content | ≤ 3 mg/mL |  | † |  |  |  | † |  |  |  |
| Vial/Package Label | Complies |  | Complies |  |  |  | Complies |  |  |  |

* Current edition of BP.

** Value given in USP.